## Instrument Parameters
Lab ID | Method Used | Method Reference | Instrument | Calibration range | Parameters
|  |  |  | (ppb, pg/ul) | m/z used for qualitative analysis | m/z used for quantitative analysis | Column Used
1 |  |  | LC/MS/MS | 1 - 50 ng/uL (10 - 500 ng/g) | 72 --> 44 (acrylamide); 75 --> 30 (d3-acrylamide) | 72 --> 55 (acrylamide); 75 --> 58 (d3-acrylamide) | Synergi 4u Hydro-RP 80A, 250 x 2.00 mm, 4 micron particles guard column: Phenomenex C18 4 x 2.00 mm
2 | No Report | No Report | No Report | No Report | No Report | No Report | No Report
3 | Internal Method |  | LC/MS/MS | 14.85 - 297 | AA: 72>27, 72>54 d3-AA: 75>30 | AA: 72>55 d3-AA: 75>58 | Shodex Rspak DE-413L polymethacrylate gel column 250 x 4.6 mm i.d., particle size 6 µm (Part. No. G307001, Showa Denko, Tokio, Japan)
4 |  |  | GC/MS/MS | 0.01 mg/L to 1 mg/L | m/z 55 and IS m/z 58 | m/z 55 and IS m/z 58 | BGB-wax 30m; 0.25 mm ID; 0.25um film
5 |  |  | GC-MS | 10 - 750 ppb | 106; 108; 152 | 150 | ZB-50 30 m; 0.25 mm I.d.; 0.25 um film thickness
6 |  | J. Rosén & K.E. Hellenäs, Analyst 2002, 127, p. 880-82 |  | 0-20 ug absolute (0-5000 ug/kg sample) | 72/55, acrylamid-d3: 75/58 | 72/54, 72/44, 72/27, | Luna 5u C18(2), 150mm*2,0 mm
7 |  |  | GC/MS (LC didn't work)
8 | Internal Method | Analytica Chimica Acta, 2004 (in press) | LC/MS/MS | 30 - 1500 | 72>72; 72>44 (AA) / 75>75; 75>44 (IS) | 72>55 (AA) / 75>58 (IS) | Merck LiChrospher® 100 CN column (250 x 4 mm I.D., 5 µm) (Darmstadt, Germany)
9 | Our method is based on the FDA's method. We exclude the spin filters and use high speed centrifuge in it place. |  | Micromass Quattro LC | 0.001 µg/mL-0.400 µg/mL | 72 - 55 72 - 44 72 - 27 | 72 - 55 75 - 58 (IS) | Phenomenex Aqua C18 5 µm 250 X 2.0 mm
10 |  | J. Agric. Food Chem. 2003, 51, 4782 | LC/MS | 50 ppb - 2 ppm | NA | 72, 73 | Waters YMC C18 AQ
11 | Based on US FDA's Method |  | LC/MS/MS | 0.5 ng/mL - 200 ng/mL (20ul), 10 - 40000 pcg on column | 72 (> 5 pcg on col.) and 44(> 40 pcg on col.) | 55 | Phenomenex Aqua 5u C18 125A, 250 x 2.00 mm
12 | Modified EPA method |  | GC/MS | approx. 0.1 microgram to 100 microgram | 150, 152 and 106 for acrylamide;155, 153, and 110 for internal std | 150 for acrylamide;155 for internal standard | DB17 (30 m)
13 | Based on US FDA's Method |  | LC/MS/MS | 1 ug/L to 50 ug/L |  | 72/55, acrylamid-d3: 75/58 | Phenomenex Synergi 4u Hydro –RP 80A, 2.0 x 250 mm with an appropriate guard column
14 | Modified | Analyst 2002, 127, 880-882 | LC/MS/MS | 0,5 - 2000 µg/kg sample | 72>55, 72>54 | 72>55, 75>58 | Hypercarb, 5µm, 50x2.1 mm
15 | No Report | No Report | No Report | No Report | No Report | No Report | No Report
16 |  |  | LC/MS/MS
17
18 | Based on US FDA's Method | see JAFC (2003) 51, 26, 7547-7554 | LC/MS/MS | 20 - 82 | 75, 72, 29, 27 | 58, 55 | Synergi 4 μ Hydro-RP 80A
19 | Based on US FDA's Method |  | LC/MS/MS | 0-1000 | 72->55, 72->27 If necessary with 75->58 and 75->29 for IS, Also, RT of sample, std, and IS's must all match | 72->55 with 75->58 for IS | C18
20 | Based on US FDA's Method |  | LC/MS/MS | 2 to 100ppb | 71.6,54.6 and 26.6 for acrylamide;74.6, 58.6 and 28.10 Internal Standard | 54.6, 58.6 | C18 Column 2x250 mm packed with 3 micron from Phenomonex

## Sample-Added Water Amount
Sample amount used and water added
Lab ID | Cereal (#1) | Peanut Butter (#2) | Chocolate (#3) | Coffee (#4) | Cereal (#5) | Water (#6) | Water Added
| g | (g) | (g) | (g) | (g) | (ml) | (ml)
1 | 1 | 1 | 1 | 1 | 1 | 1 | 10
2 | No Report | No Report | No Report | No Report | No Report | No Report | No Report
3 | 2 | 2 | 2 | 2 | 2 | 2 | 10
4 | 10 | 10 | 4 | 4 | 10 | 10 | 80
5 | 2 | 2 | 2 | 2 | 2 | 2 | 20
6 | 4 | 4 | 4 | 4 | 4 | Not Analyzed
7
8 | 2 | 2 | 2 | 2 | 2 | 2 | 50
9 | 1 | 1 | 1 | 1 | 1 | 1 | 10
10 | 6 | 6 | 6 | 6 | 6 | 6
11 | 3 | 3 | 3 | 3 | 3 | 3 | 30
12 | 2 | 2 | 2 | 0.5 | 2 | 10, 20 | 20
13 | 0.7 | 0.7 | 0.5 | 0.7 | 0.7 | 5
14 | 4 | 4 | 4 | 2 | 4 | 4 | 40
15 | No Report | No Report | No Report | No Report | No Report | No Report | No Report
16 | 4 | 4 | 4 | 4 | 4 | 10
17 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
18 | 1
19 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
20 | 1 | 1 | 1 | 1 | 1 | 1 | 9

## Sample Prep
Lab ID | Internal Standard added | When
1 | d3-acrylamide;500 ng/g | at very beginning
2 | No Report | No Report
3 | d3-acrylamide (250 µg/kg) | at very beginning
4 | 13C3-acrylamide (10 mg/L, 60 µl) | at very beginning
5 | 13C3-acrylamide | at very beginning
6 | d3-acrylamide | at very beginning
7 | n/a
8 | d3-acrylamide (200 ul, 10ug/ml) | at very beginning
9 | 13C3-acrylamide | at very beginning
10 | 13C3-acrylamide | at very beginning
11 | 13C3-acrylamide (10 ug/L, 60 µl) | at very beginning
12 | 13C-acrylamide | at very beginning
13 | 1 μg 1,2,3-13C-acrylamide | at very beginning
14 | 400 µl d3-acrylamide, 1.0 µg / ml | at very beginning
15 | No Report | No Report
16 | 13C3-acrylamide | at very beginning
17 | n/a | n/a
18 | 13C3-acrylamide | at very beginning
19 | d3-acrylamide | at very beginning
20 | 13C3-acrylamide | after extraction

## Results
| NFPA_AA_1: Cereal |  |  |  | NFPA_AA_2: PB |  |  |  | NFPA_AA-3: Chocolate |  |  |  | NFPA_AA-4: Coffee |  |  |  | NFPA_AA-5: Cereal |  |  |  | NFPA_AA-6: Water
[17 empty] | (Duplicate of NFPA_AA-1) |  |  |  | (QC; 20 ppb)
Lab ID | Result #1 | Result #2 | Ave. | Z-value | Result #1 | Result #2 | Ave. | Z-value | Result #1 | Result #2 | Ave. | Z-value | Result #1 | Result #2 | Ave. | Z-value | Result #1 | Result #2 | Ave. | Z-value | Result #1 | Result #2 | Ave. | Z-value
| (ppb) | (ppb) | (ppb) |  | (ppb) | (ppb) | (ppb) |  | (ppb) | (ppb) | (ppb) |  | (ppb) | (ppb) | (ppb) |  | (ppb) | (ppb) | (ppb) |  | (ppb) | (ppb) | (ppb)
1 | 21 | 22 | 21.5 | -0.033983 | 102 | 101 | 101.5 | -0.024741 | 222 | 211 | 216.5 | 1.9284 | 147 | 142 | 144.5 | -0.31686 | 22 | 22 | 22 | -0.24139 | 19 | 19 | 19 | -0.4246
2 | No Report | No Report |  |  | No Report | No Report |  |  | No Report | No Report |  |  | No Report | No Report |  |  | No Report | No Report |  |  | No Report | No Report
3 | 15.9 | 16.7 | 16.3 | -0.38995 | 88 | 85.2 | 86.6 | -0.22528 | 140.9 | 146.8 | 143.85 | 0.46914 | 140 | 142.7 | 141.35 | -0.3919 | 15.4 | 14.3 | 14.85 | -0.47183 | 20 | 19 | 19.5 | -0.28473
4 | 59 | 68 | 63.5 | 2.8411 | 101 | 104 | 102.5 | -0.011282 | 73 | 78 | 75.5 | -0.90377 | 191 | 168 | 179.5 | 0.51698 | 20 | 24 | 22 | -0.24139 | 27 | 19 | 23 | 0.69435
5 | 16 | 16 | 16 | -0.41049 | 86 | 84 | 85 | -0.24682 | 118 | 157 | 137.5 | 0.34159 | 142 | 149 | 145.5 | -0.29303 | 13 | 14 | 13.5 | -0.51534 | <5 | <5 | <5
6 | 46 | (60, 33) | 47 | 1.7116 | 77 | (77, 77) | 77 | -0.35449 | 48 | (38, 57) | 48 | -1.4561 | 130 | (136, 131) | 132 | -0.61465 | 30 | (23, 17) | 25 | -0.14471 | Can not do | Can not do |  | -5.7396
7 | 17 | 17 | 17 | -0.34203 | 330 | 445 | 387.5 | 3.8246 | 17 | 25 | 21 | -1.9985 | 207 | 221 | 214 | 1.3389 | 100 | 101 | 100.5 | 2.2886 | 19 | 18 | 18.5 | -0.56447
8 | <30 | <30 | <30 |  | 100 | 90 | 95 | -0.11223 | 165 | 160 | 162.5 | 0.84375 | 150 | 150 | 150 | -0.18583 | <30 | <30 | <30 |  | 17 | 17 | 17 | -0.98408
9 | <20 | <20 | <20 |  | 77 | 83 | 80 | -0.31411 | 147 | 139 | 143 | 0.45206 | 135 | 131 | 133 | -0.59083 | <20 | <20 | <20 |  | 29 | 23 | 26 | 1.5336
10 | ND (LOD<20 ppb) | ND (LOD<20 ppb) | <20 |  | 45 | 48 | 46.5 | -0.76499 | 103 | 101 | 102 | -0.37148 | 89 | 74 | 81.5 | -1.8178 | 110 | Not Enough Sample | 110 | 2.5948 | ND (LOD<20 ppb) | ND (LOD<20 ppb) | <20
11 | 12.7 | 14 | 13.35 | -0.59189 | 82.1 | 81.1 | 81.6 | -0.29258 | 122.3 | 127.1 | 124.7 | 0.084481 | 146.5 | 143.6 | 145.05 | -0.30375 | 12.4 | 13 | 12.7 | -0.54112 | 16.4 | 15.6 | 16 | -1.2638
12 | <20 (15) | <20 (15) | 15 | -0.47894 | 80 | 78 | 79 | -0.32757 | 126 | 126 | 126 | 0.11059 | 292 | 277 | 284.5 | 3.0185 | <20 (16) | <20 (15) | 15.5 | -0.45088 | <20 (18) | <20 (19) | 19.5 | -0.28473
13 | 17 | 14 | 15.5 | -0.44472 | 95 | 90 | 92.5 | -0.14587 | 144 | 153 | 148.5 | 0.56254 | 155 | 157 | 156 | -0.042883 | 14 | 13 | 13.5 | -0.51534 | 20 | 19 | 19.5 | -0.28473
14 | 14 | 14 | 14 | -0.5474 | 91.2 | 87.5 | 89.35 | -0.18827 | 140 | 152 | 146 | 0.51232 | 143 | 145 | 144 | -0.32877 | 15 | 14.5 | 14.75 | -0.47505 | 20.6 | 19.9 | 20.25 | -0.07493
15 | No Report | No Report |  |  | No Report | No Report |  |  | No Report | No Report |  |  | No Report | No Report |  |  | No Report | No Report |  |  | No Report | No Report
16 | 16.3 | 18.6 | 17.45 | -0.31123 | 97.1 | 91 | 94.05 | -0.12501 | 178 | Not Reported | 178 | 1.1551 | 158.5 | 150 | 154.25 | -0.084574 | 15.2 | 21.1 | 18.15 | -0.36548 | 23.1 | n/a | 23.1 | 0.72232
17 | n/a | n/a | 386 | 24.918 | n/a | n/a | 150 | 0.62803 | n/a | n/a | 140 | 0.3918 | n/a | n/a | 1557 | 33.3342 | n/a | n/a | 1201 | 37.7566 | n/a | n/a | ND (<50)
18 | 16 | 18 | 17 | -0.34203 | 90 | 89 | 89.5 | -0.18625 | 120 | 139 | 129.5 | 0.1809 | 148 | 139 | 143.5 | -0.34068 | 19 | 21 | 20 | -0.30585 | 18 | 18 | 18 | -0.70434
19 | 21 | 18 | 19.5 | -0.17089 | 93 | 87 | 90 | -0.17952 | 64 | 53 | 58.5 | -1.2452 | 161 | 170 | 165.5 | 0.18344 | 17 | 22 | 19.5 | -0.32197 | 32 | 26 | 29 | 2.3728
20 | 16.7 | 13 | 14.85 | -0.48921 | 75.5 | 82.8 | 79.15 | -0.32555 | 94.4 | 80.3 | 87.35 | -0.66575 | 144.3 | 192.6 | 168.45 | 0.25372 | 19.7 | 21.1 | 20.4 | -0.29296 | 18 | 19.8 | 18.9 | -0.45257
Ave. |  |  | 21.9964 |  |  |  | 103.3382 |  |  |  | 120.4941 |  |  |  | 157.8 |  |  |  | 29.49 |  |  |  | 20.5179
Std. Dev. |  |  | 14.6081 |  |  |  | 74.2991 |  |  |  | 49.7849 |  |  |  | 41.9749 |  |  |  | 31.028 |  |  |  | 3.5748
Results from Lab #17 are not included in calculations.
Results with "<" are not included in calculations.
Ave. * |  |  | 16.4542 |  |  |  | 85.5781 |  |  |  | 120.4941 |  |  |  | 154.44 |  |  |  | 18.23
Std. Dev. * |  |  | 2.2925 |  |  |  | 12.9899 |  |  |  | 49.7849 |  |  |  | 20.9529 |  |  |  | 5.461
Rel. Std. Dev. (%) * |  |  | 13.9328 |  |  |  | 15.179 |  |  |  | 41.3173 |  |  |  | 13.567 |  |  |  | 29.9562
* Note:
Results from Lab #17 are not included in calculations.
Results with "<" are not included in calculations.
| For NFPA_AA-1: |  | Results from Lab 4 and 6 are excluded.
| For NFPA_AA-2: |  | Results from Lab 7 are excluded.
| For NFPA_AA-3: |  | None of Results are excluded.
| For NFPA_AA-4: |  | Results from Lab 10 and 12 are excluded.
| For NFPA_AA-5: |  | Results from Lab 7 and 10 are excluded.
| For NFPA_AA-6: |  | None of Results are excluded.

## NFPA_AA-1
Sample ID: NFPA_AA-1
Matrix: | Cereal
Lab ID | Result #1 | Result #2 | IS Area Count (#1) | IS Area Count (#2)
| (ppb) | (ppb)
1 | 21 | 22 | 637000 | 691000
2 | No Report | No Report
3 | 15.9 | 16.7 | 54790 | 38216
4 | 59 | 68 | 4327 | 5103
5 | 16 | 16 | 10151843 | 13877435
6 | 46 | (60, 33) | 1,83*10e5 | 1,83*10e5
7 | 17 | 17 | N/A | N/A
8 | <30 | <30 | 48600 | 61200
9 | <20 | <20 | 878 | 936
10 | ND (LOD<20 ppb) | ND (LOD<20 ppb) | 87778 | 64758
11 | 12.7 | 14 | 1464786 | 1557798
12 | <20 (15) | <20 (15) | 26907 | 14228
13 | 17 | 14
14 | 14 | 14 | 203000 | 179000
15 | No Report | No Report
16 | 16.3 | 18.6 | 80303 | 130671
17 | No Report | No Report
18 | 16 | 18 | 203000 | 179000
19 | 21 | 18
20 | 16.7 | 13 | 1350 | 1510

## NFPA_AA-2
Sample ID: NFPA_AA-2
Matrix: | Peanut Butter
Lab ID | Result #1 | Result #2 | IS Area Count (#1) | IS Area Count (#2)
| (ppb) | (ppb)
1 | 102 | 101 | 283000 | 262000
2 | No Report | No Report
3 | 88 | 85.2 | 37689 | 41568
4 | 101 | 104 | 3852 | 4898
5 | 86 | 84 | 11757467 | 31387651
6 | 77 | (77, 77) | 1,95*10e5 | 1,95*10e5
7 | 330 | 445 | N/A | N/A
8 | 100 | 90 | 108000 | 146000
9 | 77 | 83 | 1116 | 1245
10 | 45 | 48 | 92426 | 113902
11 | 82.1 | 81.1 | 1839151 | 1887047
12 | 80 | 78 | 20349 | 15359
13 | 95 | 90
14 | 91.2 | 87.5 | 129000 | 123000
15 | No Report | No Report
16 | 97.1 | 91 | 136253 | 51301
17 | No Report | No Report
18 | 90 | 89 | 129000 | 123000
19 | 93 | 87
20 | 75.5 | 82.8 | 2200 | 1770

## NFPA_AA-3
Sample ID: NFPA_AA-3
Matrix: | Chocolate
Lab ID | Result #1 | Result #2 | IS Area Count (#1) | IS Area Count (#2)
| (ppb) | (ppb)
1 | 222 | 211 | 231000 | 231333.3333
2 | No Report | No Report
3 | 140.9 | 146.8 | 33282 | 32028
4 | 73 | 78 | 3216 | 3653
5 | 118 | 157 | 18240782 | 8104706
6 | 48 | (38, 57) | 1,95*10e5 | 1,95*10e5
7 | 17 | 25 | N/A | N/A
8 | 165 | 160 | 73500 | 81600
9 | 147 | 139 | 1065 | 1043
10 | 103 | 101 | 87039 | 111048
11 | 122.3 | 127.1 | 2300839 | 2165723
12 | 126 | 126 | 27242 | 24554
13 | 144 | 153
14 | 140 | 152 | 42000 | 64000
15 | No Report | No Report
16 | 178 | Not Reported | 52594 | Not reported
17 | No Report | No Report
18 | 120 | 139 | 42000 | 64000
19 | 64 | 53
20 | 94.4 | 80.3 | 1450 | 1950

## NFPA_AA-4
Sample ID: NFPA_AA-4
Matrix: | Coffee
Lab ID | Result #1 | Result #2 | IS Area Count (#1) | IS Area Count (#2)
| (ppb) | (ppb)
1 | 147 | 142 | 146500 | 145500
2 | No Report | No Report
3 | 140 | 142.7 | 31279 | 25663
4 | 191 | 168 | 1300 | 2391
5 | 142 | 149 | 11131159 | 6912101
6 | 130 | (136, 131) | 1,95*10e5 | 1,95*10e5
7 | 207 | 221 | N/A | N/A
8 | 150 | 150 | 41400 | 32200
9 | 135 | 131 | 1347 | 1377
10 | 89 | 74 | 61974 | 57988
11 | 146.5 | 143.6 | 2186346 | 2145319
12 | 292 | 277 | 34484 | 25502
13 | 155 | 157
14 | 143 | 145 | 25000 | 24000
15 | No Report | No Report
16 | 158.5 | 150 | 66000 | 30458
17 | No Report | No Report
18 | 148 | 139 | 25000 | 24000
19 | 161 | 170
20 | 144.3 | 192.6 | 1390 | 1070

## NFPA_AA-5
Sample ID: NFPA_AA-5
Matrix: | Cereal
Lab ID | Result #1 | Result #2 | IS Area Count (#1) | IS Area Count (#2)
| (ppb) | (ppb)
1 | 22 | 22 | 718500 | 705000
2 | No Report | No Report
3 | 15.4 | 14.3 | 35071 | 31472
4 | 20 | 24 | 3534 | 3321
5 | 13 | 14 | 8849325 | 21065714
6 | 30 | (23, 17) | 1,83*10e5 | 1,83*10e5
7 | 100 | 101 | N/A | N/A
8 | <30 | <30 | 68500 | 57500
9 | <20 | <20 | 796 | 787
10 | 110 | Not Enough Sample | 68881 | N/A
11 | 12.4 | 13 | 1758256 | 1774208
12 | <20 (16) | <20 (15) | 34812 | 22788
13 | 14 | 13
14 | 15 | 14.5 | 179000 | 140000
15 | No Report | No Report
16 | 15.2 | 21.1 | 84701 | 111495
17 | No Report | No Report
18 | 19 | 21 | 179000 | 140000
19 | 17 | 22
20 | 19.7 | 21.1 | 1260 | 1100

## NFPA_AA-6
Sample ID: NFPA_AA-6
Matrix: | 20 ppb in water
Lab ID | Result #1 | Result #2 | IS Area Count (#1) | IS Area Count (#2)
| (ppb) | (ppb)
1 | 19 | 19 | 1640000 | 1660000
2 | No Report | No Report
3 | 20 | 19 | 31411 | 44936
4 | 27 | 19 | 4517 | 5121
5 | <5 | <5 | 6044399 | 7760382
6 | Can not do | Can not do | N/A | N/A
7 | 19 | 18 | n/a | n/a
8 | 17 | 17 | 220000 (0.5 µg IS added) | 7170 (0.02 µg IS added)
9 | 29 | 23 | 997 | 1031
10 | ND (LOD<20 ppb) | ND (LOD<20 ppb) | 137399 | 110644
11 | 16.4 | 15.6 | 1838923 | 1820561
12 | <20 (18) | <20 (19) | 2138 | 10007
13 | 20 | 19 | 3714482 | 4411696 | Revised on 4/12/04 | Process Blank
14 | 20.6 | 19.9 | 279000 | 673000 |  | IS area counts: 4849969
15 | No Report | No Report
16 | 23.1 | n/a | 289397 | n/a
17 | No Report | No Report
18 | 18 | 18 | 279000 | 673000
19 | 32 | 26
20 | 18 | 19.8 | 1680 | 2140